Procedure: Point of Care Urine Pregnancy – Quidel QuickVue One-Step Urine hCG

| Prepared by |  | Date Adopted |  | Supersedes Procedure # |
| --- | --- | --- | --- | --- |
| Jeronimo Belgrave | April 28, 2003 |  | Abbott Test Pack Urine HCG |  |

| Review Date |  | Revision Date |  |  |  | Signature |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Distributed to |  |  | # of Copies | Distributed to |  |  | # of Copies |
| EMMERGENCY DEPARTMENT |  |  | 1 |  |  |  |  |
| MIDWOOD SATELLITE |  |  | 1 |  |  |  |  |
| LEFFERTS SATELLITE |  |  | 1 |  |  |  |  |
| THROOP SATELLITE |  |  | 1 |  |  |  |  |

PRINCIPLE:  The Quick Vue One-Step hCG Urine Test is a sensitive immunoassay for the qualitative detection of human chorionic gonadotropin (HCG) in urine for the early detection of pregnancy.  Urine is added to the Sample Well on the Test cassette.  If hCG is present  in the specimen at a level of 25 mIU/ml or greater, a pink-to-purple Test (T) line will appear along with a blue Control (C) line in the result Window.  If hCG is present at lower levels, or not present in the specimen, only a blue Control Line will appear in the result window.

SPECIMEN:  Urine collected in a clean container.  First morning specimens generally contain the highest concentrations of hCG and are recommended for early detection of pregnancy.  However, any urine specimen is suitable for testing.

Patient Preparation:  No patient  preparation is necessary.

Handling Conditions: A freshly voided and appropriately labeled urine specimen may be kept at room temperature for 8 hours or stored at 2-8oC for up to three days.  Do not freeze specimens.

EQUIPMENT AND MATERIALS:

Equipment:  N/A

Materials: 75 individually wrapped test Cassettes.

*Each cassette contains murine monoclonal antibody and caprine

polyclonal antibody to hCG.

75 disposable pipets

# Preparation:  N/A

Storage Requirements:  Store kit at room temperature 59 - 86oF  (15 - 30oC), out of direct sunlight.  Kit contents are stable until the expiration date printed on the outer box carton.

LEVEL OF FUNCTION:  Unit staff in collaboration with Laboratory staff.

LEVEL OF PERSONNEL:  Physicians

Physicians Assistants

Registered Nurses

Laboratory Personnel

CALIBRATION:  N/A

QUALITY CONTROL:  The Quick Vue test contains built-in control features.  The development of the blue procedural Control Line next to the letter “C” is a Positive procedureal control.  If this line does not develop, the test is invalid.

EXTERNAL CONTROL:  Hycor Liquitrol Normal and Abnormal (Positive and Negative) controls will be tested and documented once per day or  on day of testing and initialed by the Nurse performing QC.  Control documentation record will be submitted to the Point of Care Coordinator at the end of each month for review.

PROCEDURE - STEPWISE:

- Remove the QuickVue Test Cassette from the foil pouch and place it on a clean, dry surface.

- Using one of the disposable pipets supplied, add 3 DROPS of urine to the Round Sample Well on the Test Cassette.

- Wait three minutes and read.

Note:  Must be read at THREE  minutes.

RESULTS:

POSITIVE:  The appearance of any pink-to-purple line next to the letter “T” in the Result Window, along with a blue procedural Control Line next to the letter “C”.

NEGATIVE:  The appearance of the blue procedural Control Line next to the letter “C”  only and no pink-to-purple Test Line next to the letter “T”.

NO RESULT:  If no blue procedural Control Line appears, the test result is invalid and the specimen must be retested.

REPORTING RESULTS:

All patient results must be documented in a Patient Logbook containing the following information:

DATE, TIME, PATIENT’S NAME, PATIENT’S MR#, ORDERING PHYSICIAN, TEST RESULT, NURSE.

Upon completion, each patient’s urine hCG result will be entered into Cerner LIS  via one of the following methods:

a)	Fill out patient result sheet containing Name of Patient, MR# and Test Result.  Result forms are kept in the Patient Log Binder.  Give completed Result Form to Clerk who will order and result in Cerner.

b)	Order Point of Care Test in Cerner:

1)  Click on Department Order Entry

- Enter Patient’s Name in Person Name Field

- Type “POC” in Orderable Field

- Type in result (Pos. or Neg.)

- Choose “Central Receiving” in Specimen Received Location Field.

- Click on Submit Order

LIMITATIONS OF THE PROCEDURE:  The contents of this kit are for use in the

qualitative detection of hCG in urine only.  Sensitivity – 25mIU/ml.

REFERENCES:   1)  Quidel QuickVue Test Pack Insert.