SUNY DOWNSTATE MEDICAL CENTER

UNIVERSITY HOSPITAL OF BROOKLYN

POLICY AND PROCEDURE

No:  Section 15K

Subject:    PROCEDURE FOR CLEANING AND

DISINFECTION OF FLEXIBLE FIBEROPTIC 			Page:  1 of  2

ENDOSCOPES (FFE) USED IN G.I. ENDOSCOPY										Original Issue date: 1980

Prepared by:  George Allen, Ph.D., CIC			Supercedes: August 2006

Reviewed/Approved by:					Approval Date: August 2007

Michael Augenbraun, MD, FACP

Distribution:	Infection Control Manual

Issued by:  Infection Control Program

___________________________________________________________________________

I:	Purpose: To provide guidelines for effectively reprocessing of endoscopes.

II:	Definition: None

III:	Policy: The procedures outlined will be used when reprocessing endoscopes.

IV:	Role Responsibility: Department head/designee will ensure compliance

V:	Procedures/Guidelines:

Scrupulous mechanical cleaning of the insertion tube and channels, using an antimicrobial

solution is carried out.  It must be done immediately after use to prevent the drying of

secretions, to remove organic material and lower the number of microorganisms.

To facilitate the removal of proteinaceous material, an enzymatic cleaner may be used

following the recommendations from the manufacturer of the scope.

Inspection of the equipment for damage is carried out at all stages of handling.  A "leak test" is

performed to check for punctures. The entire outside of the scope is wiped with a germicide

detergent. High level disinfection of the endoscopic insertion tube and all channels is performed

with (OPA) ortho-phtalaldehyde (EPA registered) by aspiration into the channel and then

soaking in OPA for a minimum 20 minutes. Adequate rinsing utilizing sterile water must follow.

This will prevent possible residual toxic effects of the disinfectant.  Since tap water may contain

microbes, including Pseudomonas spp. and Mycobacteria spp.; rinsing must be done with

sterile water. The insertion tube and inner channels shall be thoroughly and immediately air

dried after cleaning and prior to storage as it is known that bacteria will multiply in a moist

environment. 70% alcohol may be used as a final rinse before storing.

If Sterrad TM or plasma is used, it is imperative that meticulous cleaning be accomplished as

described in 1 above.

Because of the spring-like structural configuration reusable, accessories such as biopsy forceps

and cytology brushes are extremely difficult to clean and disinfect.  After immediate surface

cleaning with a germicide detergent solution,  use either steam under pressure, gas or plasma.

These methods have the capability of penetrating the spring-like structures.

PROCEDURE FOR THE CLEANING & DISINFECTION OF FIBEROPTIC BRONCHOSCOPE

Determine and document the efficacy of the OPA utilizing the test strips and the log.

Fill the disinfecting tray with activated OPA.

Clean the endoscopes thoroughly, following the manufacturer's directions, before disinfecting.

Connect the syringe to the biopsy port. Immerse the insertion tube into the OPA solution.  Do

not immerse the control unit or the light guide connector.  Secure the control unit on the lip of

the tray with the safety strap provided.

Draw 3cc of OPA solution through the channel into the syringe.  Allow the OPA solution to

contact the insertion tube and channel for a minimum of 20 minutes. Expel the OPA solution

from the channel by depressing the syringe and remove the insertion tube from the OPA

solution.

Rinse the insertion tube and the channel (using the syringe) with sterile water.  Rinse at least

four times  using fresh sterile water for each rinse.

Wipe off the remainder of the instrument with an alcohol dampened 4"x4" gauze pad.

Force air through the channel to dry the instrument thoroughly.  This must be done before

storing.  A final rinse with alcohol before storage is also acceptable.

Because of the spring-like structural configuration, reusable accessories such as biopsy

forceps and cytology brushes are extremely difficult to clean and disinfect.  After immediate

surface cleaning with a germicide detergent solution and rinsing, use either steam under

pressure, or plasma sterilization.  These methods have the capability of penetrating the spring-

like structures.

To ensure potency of the OPA solution, test strips must be used to monitor the solution.

Personnel must utilize the applicable personal protective equipment, including gloves, gowns

and face protective devices.  To reduce the dispersal of OPA vapors, containers must be tightly

covered while the equipment is being soaked.

SAMPLING PROCEDURE

Whenever transmission of infection is suspected, random sampling of endoscopes may be initiated utilizing the following procedures: Collect supplies: Culture medium (broth) from Microbiology: mucous trap; sterile gloves, bowl and saline. Using aseptic technique, suction the culture medium through the scope into the mucous trap. Detach and cover the mucous trap, label and send the specimen to Microbiology. Document results in a log book.

VI:	Reason for Revision: Review

VII:	Attachments:

VIII:	References: ASGE, APIC, CDC

| Date Reviewed | Revision Required (Check One) |  | Responsible Staff Name and Title |
| --- | --- | --- | --- |
| 8/15/2006 | Yes | Yes | George Allen, Ph.D. Dir. of Infection Control |
| 8/07 | Yes | Yes | G Allen, PhD, CIC, CNOR |
|  | Yes | Yes |  |
|  | Yes | Yes |  |